## Slide 1: METHODS OF MEASURING PROCESS VARIABLES AND THEIR CONTROL

## Slide 2: Temperature
The temperature in a vessel or pipe is one of the most important parameters to monitor and control in any process.
Measurement
It is measured by
  mercury-in-glass thermometers,
  bimetallic thermometers,
  pressure bulb thermometers,
  thermocouples,
  metal-resistance thermometers or thermistors.
Metal-resistance thermometers and thermistors are used in most fermentation applications.
Control
  The use of water jackets or pipe coils within a fermenter as a means of temperature control.
  and a cooling water supply; these are on or off depending on the need for heating or cooling.
  In large fermenters, where heating during the fermentation is not normally required, a regulatory valve at the cooling-water inlet may be sufficient to control the temperature.
  Low agitation speeds are often essential in animal cell culture vessels to minimize shear damage.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 3: Flow of fluids
Flow measurement and control of both gases and liquids is important in process management .
Measurement
  for measuring gas flow to a fermenter is by means of a variable area meter.
  The most commonly used example is a rotameter
  Rotameters can also be used to measure liquid flow rates, provided that abrasive particles or fibrous matter are not present.
  The use of oxygen and carbon dioxide gas analysers for effluent gas analysis requires the provision of very accurate gas-flow measurement if the analysers are to be used effectively.
Control
  Control of gas flow is usually by needle valves.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 4
The mass flow rate of the gas, Q, can be calculated from the specific heat equation:
H = Q Cp(T2 – T1)
Q = H / Cp(T2 – T1)
Where,
H= heat transferred,
Q= mass flow rate of the gas,
Cp = specific heat of the gas,
T1= temperature of gas before heat is transferred to it,
T2 = temperature of gas after heat is transferred to it.
In batch and fed-batch culture fermenters, a cheaper alternative is to measure flow rates indirectly by load cells.
Load cells also monitor the increases and decreases in weight of the various vessels at regular time intervals.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 5: Pressure
Measurement
  Pressure is one of the crucial measurements.
  Is most important for safety purpose.
  The measurement of pressure is also important in media sterilization.
  In a fermenter, pressure will influence the solubility of gases and
  Contribute to the maintenance of sterility when a positive pressure is present.
  One of the standard pressure measuring sensors is the Bourdon tube pressure gauge
  Diaphragm gauge can be used when a vessel or pipe is to be operated under aseptic conditions.
Control
During normal operation a positive head pressure of 1.2 atmospheres (161 kN- 1) absolute is maintained in a fermenter to assist in the maintenance of aseptic conditions.
The correct pressure in different components is
  maintained by regulatory valves or safety valves and
  controlled by associated pressure gauges.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 6
FIG. 1. 'c' Bourdon tube pressure gauge (Liptak, 1969).
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 7
FIG. 2. Nested diaphragm-type pressure sensor (Liptak, 1969).
Changes in pressure cause movements of the diaphragm capsule which are monitored by a mechanically levered pointer.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 8: Foam sensing and control
Foaming is a problem in most of the microbiological processes, due to a component in the medium or some factor produced by the micro-organism.
The most common cause of foaming is due to proteins in the medium, such as corn-steep liquor, Pharmamedia, meal, soybean meal, yeast extract or meat extract (Schugerl, 1985).
The foaming can cause removal of cells from the medium which will lead to autolysis and the further release of microbial cell proteins will probably increase the stability of the foam.
It is common practice to add an antifoam to a fermenter when the culture starts foaming above a certain predetermined level.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 9
Control by
  Addition of chemical antifoams which are surface active agents, breaking the foam by reducing the surface tension at broth and destabilizing protein films.
  A number of mechanical antifoam devices attached to the agitator shaft above the surface of the broth such as-
    Discs foam breaker,
    propellers,
    brushes or hollow cones
  The foam is broken down when it is thrown against the walls of the fermenter.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 10
A foam sensing and control unit is shown in Fig. 3.
A probe is inserted through the top plate of the fermenter.
Normally the probe is a stainless-steel rod, which is insulated except at the tip, and set at a defined level above the broth surface.
When the foam rises and touches the probe tip, a current is passed through the circuit of the probe, with the foam acting as an electrolyte and the vessel acting as an earth.
The current actuates a pump or valve and antifoam is released into the fermenter for a few seconds which result into breaking and reduction of foam formation.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 11
FIG. Foam sensing and control unit.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 12
Antifoams are surface active agents, reducing the surface tension in the foams and destabilizing protein films by
  (a) hydrophobic bridges between two surfaces,
  (b) displacement of the absorbed protein, and
  (c) rapid spreading on the surface of the film (Van't Riet and Van Sonsbeck, 1992).
ANTIFOAMS
Foaming is a problem in most of the microbiological processes, due to a component in the medium or some factor produced by the micro-organism.
The most common cause of foaming is due to proteins in the medium, such as corn-steep liquor, Pharmamedia, meal, soybean meal, yeast extract or meat extract (Schugerl, 1985).
The foaming can cause removal of cells from the medium which will lead to autolysis and the further release of microbial cell proteins will probably increase the stability of the foam.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 13: The following compounds are used as antifoam and meet most of these properties have been found to be most suitable in different fermentation processes (Solomons, 1969;Ghildyal et al., 1988):
Alcohols; stearyl and octyl decanol.
2. Esters.
3. Fatty acids and derivatives, particularly glycerides, which include cottonseed oil, linseed oil, say-bean oil, olive oil, castor oil, sunflower oil, rapeseed oil and cod liver oil.
4. Silicones.
5. Sulphonates.
6. Miscellaneous; Alkaterge C, oxazaline, poly-propylene glycol.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 14: An ideal antifoam should have the following properties:
1.Should disperse readily and have fast action on an existing foam.
2. Should be active at low concentrations.
3. Should be long acting in preventing new foam formation.
4. Should not be metabolized by the microorganism.
5. Should be non-toxic to the micro-organism.
6. Should be non-toxic to humans and animals.
7. Should not cause any problems in the extraction and purification of the product.
8. Should not cause any handling hazards.
9. Should be cheap.
10. Should have no effect on oxygen transfer.
11. Should be heat sterilizable.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 15
Measurement
  In most aerobic fermentations it is essential to ensure that the dissolved oxygen concentration does not fall below a specified minimal level.
  Monitored by sterilizable oxygen electrodes.
  These electrodes measure the partial pressure of the dissolved oxygen and not the dissolved oxygen concentration.
  Therefore at equilibrium, the probe signal of an electrode will be determined by:
  P(O2) = C(O2) x PT
where
P(O2) is the partial pressure of dissolved oxygen sensed by the probe,
C(O2) is the volume or mole fraction of oxygen in the gas phase,
PT is the total pressure.
The actual reading is normally expressed as percentage saturation with air at atmospheric pressure,
100% dissolved oxygen means a partial pressure of approximately 160 mmHg.
Measurement and control of dissolved oxygen
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 16: Agitator and Shaft Power
Measurement and Control
A variety of sensors can be used to measure the power consumption of a fermenter.
On a large scale, a watt meter attached to the agitator motor will give a fairly good indication of power uptake.
Torsion dynamometers can be used in small-scale applications.
Rate of stirring
  In all fermenters it is important to monitor the rate of rotation (in rpm) of the stirrer shaft.
  Measured by the used of tachometer.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 17: Weight and Microbial biomass
Measurement by load cell .
A load cell offers a convenient method of determining the weight of a fermenter or feed vessel.
A load cell is essentially an elastic body, usually a solid or tubular steel cylinder.
Changes of resistance with strain which are proportional to load are determined by appropriate electrical apparatus.
The change in weight in a known time interval can be used indirectly as a measure of liquid flow rates.
Load cells also used to monitor
  feed rates from medium reservoirs,
  acid and base utilization for pH control and
  the use of antifoam for foam control.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 18: pH measurement and control
In batch culture the pH of an actively growing culture will not remain constant for long time.
Rapid changes in pH can often be reduced by
  the careful design of media, particularly in the choice of carbon and nitrogen sources, and
  the incorporation of buffers or
  batch feeding.
The pH may be further controlled by
  the addition of appropriate quantities of ammonia or sodium hydroxide if broth become too acidic, or
  the addition of appropriate quantities of sulphuric acid if broth become too alkaline.
pH measurement is now routinely carried out
  using a combined glass reference electrode that will
  withstand repeated sterilization at temperatures of 1210C and pressures of 138 kN m- 2.
  The electrodes may be silver/silver chloride with potassium chloride or special formulations (e.g. Friscolyt by Ingold) as an electrolyte.
  Occasionally calomel/mercury electrodes are used.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 19
The measurement of dissolved carbon dioxide is possible with an electrode, since a pH or voltage change can be detected as the gas goes into solution.
Measured by electrode consisted of a combined pH electrode with a bicarbonate buffer (pH 5) surrounding the bulb and ceramic plug, with the solution.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition

## Slide 20: Redox
It is a measure of the oxidation-reduction potential of a biological system and
Can be determined as a voltage (mV), the value in any system depending on the equilibrium of:
Reduced form  Oxidized form + electron (s)
(negative value) (positive value)
The measuring electrode consists of gold, platinum or iridium which is welded to a copper lead.
Notes: Principles of Fermentation Technology,- Peter F. Stanbury, Allen Whitaker, Stephen J. Hall, Second Edition